## Slide 1: Data Augmentation for Epileptic Seizure Detection based on 3D acceleration, heart rate and temperature
1
First examiner: Prof. X. Cologne University of Applied Sciences
Second examiner: Y, Fraunhofer Institute for Software and Systems Engineering
Colloquium for Master’s Thesis by Maleyka Seyidova
Cologne University of Applied Sciences
in cooperation with
Fraunhofer Institute for Software and Systems Engineering

## Slide 2: Content
Motivation and Objective
Dataset
Method & Implementation
Results & discussion
Conclusion
Outlook

## Slide 3: Content
Motivation and Objective
Dataset
Method & Implementation
Results & discussion
Conclusion
Outlook

## Slide 4: Motivation
Epilepsy
  characterized by recurrent seizures
  can lead to serious injuries and life-threatening situations
Source: WATSON, JL, and Lee Memorial. “Epilepsy Monitoring Unit Helps Physicians Make Diagnosis While Patients Sleep.” Fort Myers Florida Weekly -, Fort Myers Florida Weekly, 10 Sept. 2014, fortmyers.floridaweekly.com/articles/epilepsy-monitoring-unit-helps-physicians-make-diagnosis-while-patients-sleep/. Accessed 30 May 2023.
‌
Automated seizure detection systems
  documentation & alarm system
  requires a large amount of data
Unpredictable nature of seizures
  Highly imbalanced dataset

## Slide 5
MOND based on the EpiTect
Mobile, smart neurosensory system for the detection and documentation of epileptic seizures in everyday life
3D acceleration
temperature
heart rate - PPG
Source: Houta, Salima, Pinar Bisgin, and Pascal Dulich. "Machine learning methods for detection of Epileptic seizures with long-term wearable devices." Eleventh International Conference on eHealth, Telemedicine, and Social Medicine. 2019.

## Slide 6: Objective
Data Augmentation
  Increase the amount of available data for MOND – Generate synthetic seizures
Seizure sensitivity
False Alarm Rate
Standard Time Series Augmentation
Generative Adversarial Networks
  Improve the performance of available seizure detection system
    Motoric seizure detection

## Slide 7: Content
Motivation and Objective
Dataset
Method & Implementation
Results & discussion
Conclusion
Outlook

## Slide 8: Dataset
3D acceleration (x-axis, y-axis, z-axis)
  Sampling frequency: 50Hz
  Unit: g

## Slide 9: Dataset
Heart rate
  Sampling frequency: 1Hz
  Unit: bpm
Temperature
  Sampling frequency: 1Hz
  Unit: °C

## Slide 10: Dataset
Two datasets
  Three-feature dataset: only 3D-acceleration
  Five-feature dataset: 3D-acceleration, Heart rate and Temperature
Dataset - motoric seizures
  FIAS – Focal Impaired Awareness Seizure
  FAS – Focal Aware Seizure
  TCS – Tonic-Clonic Seizure
  GMS – Generalized Motor Seizure
Total 109 motoric seizures corresponding to 46 patients

## Slide 11: Content
Motivation and Objective
Dataset
Method & Implementation
Results & discussion
Conclusion
Outlook

## Slide 12: Method & Implementation
Training CNN with Augmented Data
Seizure detection without Data Augmentation
Standard Time Series Data Augmentation
Data Preprocessing
Train-Test Split
CNN Architecture
Three-inputs CNN
Five-inputs CNN
Evaluation
Selection of Baseline
Workflow
Techniques Implementation
Single Techniques
Combination
GAN- based Time Series Data Augmentation
EpilepsyGAN
Workflow
TTS-GAN
Workflow
Training
Training
Similarity Evaluation
Metrics Monitoring
Windowing
Labeling
Seizure Detection with Data Augmentation

## Slide 13: Method & Implementation
Standard Time Series Data Augmentation
Data Preprocessing
Train-Test Split
CNN Architecture
Three-inputs CNN
Five-inputs CNN
Evaluation
Selection of Baseline
Workflow
Techniques Implementation
Single Techniques
Combination
GAN- based Time Series Data Augmentation
EpilepsyGAN
Workflow
TTS-GAN
Workflow
Training
Training
Similarity Evaluation
Metrics Monitoring
Training CNN with Augmented Data
Windowing
Labeling
Seizure Detection with Data Augmentation
Seizure detection without Data Augmentation

## Slide 14
Epileptic Patients Dataset
Windowing Interpolation
CNN Training
Undersampling Standardization
Five-feature dataset
acc x, acc y, acc z,
heart rate, temperature
Three-feature dataset
acc x, acc y, acc z
Data preprocessing

## Slide 15
If a window fails in any single feature, all the corresponding windows will also be discarded
Total expected values inside the window
  3D acceleration: 500 values
  Heart rate & Temperature: 10 values
Interpolation threshold: 15%
No overlap
Data preprocessing - Windowing
10 s window

## Slide 16
Dataset | # of Seizure windows | # of Non-seizure windows
Three- feature | 941 | 648399
Five-feature | 785 | 584737
Labelling the windows:
  if any part of the window lies within the seizure area, the window is labeled as “1:seizure”
  else, the window is labeled as “0:non-seizure”
Three-feature dataset | 
Seizure Types | # of events
FIAS | 60
FAS | 21
TCS | 11
GMS | 2
Five-feature dataset | 
Seizure Types | # of events
FIAS | 54
FAS | 21
TCS | 11
GMS | 2
total: 1803.72 h
seizure: 2.61 h
total: 1626.45 h
seizure: 2.18 h
Data preprocessing - Labelling

## Slide 17
Three-feature dataset |  | 
Seizure Types | # of events Train | # of events Test
FIAS | 48 | 12
FAS | 18 | 3
TCS | 7 | 4
GMS | 1 | 1
Five-feature dataset |  | 
Seizure Types | # of events Train | # of events Test
FIAS | 42 | 12
FAS | 18 | 3
TCS | 7 | 4
GMS | 1 | 1
Train-Test split is done based on the Patients:
  Train: 37 patients Test: 9 patients
Test set:
  all motoric seizure types
  patients with high and low numbers of seizures
Data preprocessing – Train-Test split

## Slide 18
Dataset | # of Seizure windows | # of Non-seizure windows
Three- feature | 791 | 581295
Five-feature | 660 | 525367
Dataset | # of Seizure windows | # of Non-seizure Windows
Three- feature | 150 | 67104
Five-feature | 125 | 59370
Train
Test
Undersampling is applied to the train set to balance the classes
Z-standardization is applied to both train and test set
Three-feature dataset |  | 
Seizure Types | # of events Train | # of events Test
FIAS | 48 | 12
FAS | 18 | 3
TCS | 7 | 4
GMS | 1 | 1
Five-feature dataset |  | 
Seizure Types | # of events Train | # of events Test
FIAS | 42 | 12
FAS | 18 | 3
TCS | 7 | 4
GMS | 1 | 1
Train-Test split is done based on the Patients
  Train: 37 patients Test: 9 patients
Test set:
  all motoric seizure types
  patients with high and low numbers of seizures
Data preprocessing – Train-Test split

## Slide 19
Conv1D, 2, 64,ReLU
Max-Pool, 2
Dropout, 0.7
Global Avg. Pool
Concatenation
Dense, 50, ReLU
Dropout, 0.7
Dense, 2, SoftMax
Conv1D, 2, 64, ReLU
Max-Pool, 2
Dropout, 0.7
Global Avg. Pool
Dense, 50, ReLU
Dropout, 0.7
Conv1D, 2, 64, ReLU
Max-Pool, 2
Dropout, 0.7
Global Avg. Pool
Dense, 50 , ReLU
Dropout, 0.7
Seizure (1) / Non-Seizure (0)
Acc x
Acc y
Acc z
CNN architecture – CNN with three inputs

## Slide 20
Conv1D, 2, 64,ReLU
Max-Pool, 2
Dropout, 0.7
Global Avg. Pool
Concatenation
Dense, 50, ReLU
Dropout, 0.7
Dense, 2, SoftMax
Conv1D, 2, 64, ReLU
Max-Pool, 2
Dropout, 0.7
Global Avg. Pool
Dense, 50, ReLU
Dropout, 0.7
Conv1D, 2, 64, ReLU
Max-Pool, 2
Dropout, 0.7
Global Avg. Pool
Dense, 50 , ReLU
Dropout, 0.7
Seizure (1) / Non-Seizure (0)
Conv1D, 2, 64, ReLU
Max-Pool, 2
Dropout, 0.7
Global Avg. Pool
Dense, 50 , ReLU
Dropout, 0.7
Conv1D, 2, 64, ReLU
Max-Pool, 2
Dropout, 0.7
Global Avg. Pool
Dense, 50 , ReLU
Dropout, 0.7
Acc x
Acc y
Acc z
Heart Rate
Temperature
CNN architecture – CNN with five inputs
20

## Slide 21
21
Threshold Moving
Choice of Loss Function
Binary Cross Entropy
Binary Focal Cross Entropy
CNN Model
G-Mean
Hyperparameter Tuning
Hyper Band Algorithm
Random Search
Sampling
Under-Sampling with different ratios
Baseline
selection
Additional
experiments
Selection of Baseline Model

## Slide 22
  Seizure Sensitivity – over the whole seizure interval
  Window Sensitivity – over individual window
  False Alarms per 24 hours (FA/24h)
  Specificity
22
1
1
1
0
0
0
0
0
0
0
0
Seizure 1
Detected
Not Detected
1
0
0
0
0
0
0
1
0
0
Detected
Seizure 2
Seizure 3
Evaluation

## Slide 23
Standard Time Series Data Augmentation
Data Preprocessing
Train-Test Split
CNN Architecture
Three-inputs CNN
Five-inputs CNN
Evaluation
Selection of Baseline
Workflow
Techniques Implementation
Single Techniques
Combination
Windowing
Labeling
Seizure Detection with Data Augmentation
GAN- based Time Series Data Augmentation
EpilepsyGAN
Workflow
TTS-GAN
Workflow
Training
Training
Similarity Evaluation
Metrics Monitoring
Training CNN with Augmented Data
Seizure detection without Data Augmentation

## Slide 24
Time Series Data Augmentation
Generative Adversarial Networks (GANs)
TTS-GAN
Standard Data Augmentation Techniques
Jittering
Rotation
Permutation
Time Warping
Magnitude Warping
Window Slicing
Combination of Techniques
Rotation + Time Warping
Rotation + Permutation
Permutation + Time Warping
Rotation + Permutation + Time Warping
EpilepsyGAN
24
Time Series Data Augmentation

## Slide 25
Standard Time Series Data Augmentation
Data Preprocessing
Train-Test Split
CNN Architecture
Three-inputs CNN
Five-inputs CNN
Evaluation
Selection of Baseline
Workflow
Techniques Implementation
Single Techniques
Combination
Windowing
Labeling
Seizure Detection with Data Augmentation
GAN- based Time Series Data Augmentation
EpilepsyGAN
Workflow
TTS-GAN
Workflow
Training
Training
Similarity Evaluation
Metrics Monitoring
Training CNN with Augmented Data
Seizure detection without Data Augmentation

## Slide 26
Epileptic Patients Dataset
Extraction of Seizure Events
Applying Augmentation
Synthetic Seizures
10 sec Windowing
CNN Training
10sec
Windowing
Original data
Synthetic data
Varying sampling frequency of acceleration vs. heart rate, temperature
Some techniques are applied only to acceleration data while keeping the heart rate and temperature original
Some techniques are applied to all features: acceleration data, heart rate, and temperature
Standard Time Series Data Augmentation - Workflow

## Slide 27
Augmentation with Jittering
Jittering – addition of noise
Std = 0.02 ; 0.05

## Slide 28
Augmentation with Rotation

## Slide 29
Augmentation with Permutation
Permutation – segmenting the seizure interval and rearranging these segments
Number of segments n=2 ; n=3; n=5

## Slide 30
Augmentation with Rotation + Permutation

## Slide 31
Augmentation with Permutation + Time Warping

## Slide 32
Synthetic samples from each single technique

## Slide 33
Synthetic samples from combination of techniques

## Slide 34
Standard Time Series Data Augmentation
Workflow
Techniques Implementation
Single Techniques
Combination
GAN- based Time Series Data Augmentation
EpilepsyGAN
Workflow
TTS-GAN
Workflow
Training
Training
Similarity Evaluation
Metrics Monitoring
Training CNN with Augmented Data
Data Preprocessing
Train-Test Split
CNN Architecture
Three-inputs CNN
Five-inputs CNN
Evaluation
Selection of Baseline
Windowing
Labeling
Seizure detection without Data Augmentation
Seizure Detection with Data Augmentation

## Slide 35
GANS – Workflow
Epileptic Patients dataset
Extraction of seizure Events
Windowing 10s
0% , 50% , 70% overlap
Training
TTS-GAN
Synthetic Seizures data
Training
EpilepsyGAN
Windowing 10s
0% overlap
Extraction of non-seizure Events
Synthetic Seizures data
Synthetic Seizure data = Synthetic 3-D acceleration (x, y, z) seizure samples

## Slide 36
Overlapping | # of Seizure windows | # of Seizure & Non-seizure pairs
0% | 525 | 5250
50% | 1009 | 10090
70% | 1484 | 14840
Overlapping | # of Seizure windows
0% | 585
50% | 1113
70% | 1638
GANS – Training
TTS – GAN
EpilepsyGAN

## Slide 37
Dynamic Time Warping (DTW)
  Measures the similarity between two-time series signals by finding the optimal alignment of the two signals such that the difference between them is minimized
Cosine Similarity
  Calculated between the feature vectors of two time series and measures the similarity by calculating the cosine of the angle between them
Mean Squared Error
  Assesses the difference between the real and synthetic values of a time series at each time point. The lower MSE value refers to the higher similarity between the two-time series
PCA and t-SNE
  Visualizing the distribution of higher dimensional data in a two-dimensional space
Similarity Evaluation

## Slide 38
TTS-GAN – Metrics monitoring
Synthetic 1
Synthetic N
. . .
dtw_1
dtw_2
dtw_N
dtw_1
dtw_2
dtw_N
. . .
. . .
All real
All real
Minimum dtw
Minimum dtw
. . .
Final DTW value
Average
Best similarity match between synthetic and real seizures
  single synthetic is compared with all real
  best match is found – min. DTW, MSE; max. Cos sim
  average of all best matches
38
Monitoring at each epoch
Saving the best model
  Lowest MSE value
  Lowest DTW value
  Highest Cosine similarity

## Slide 39
Standard Augmentation Techniques |  | 
Proportion | # of Seizure windows Five-feature dataset | # of Seizure windows Three-feature dataset
1x | 660 + 490 | 791 + 540
2x | 660 + 980 | 791 + 1080
4x | 660 + 1960 | 791 + 2160
8x | 660 + 3912 | 791 + 4320
Training CNN with augmented data
Synthetic data is added to the original training set
Undersampling: the same number of seizures & non-seizures
CNN network is trained with augmented data
GANs | 
Proportion | # of Seizure windows Three-feature dataset
10% | 791 + 80
30% | 791 + 238
50% | 791 + 396
80% | 791 + 633

## Slide 40: Content
Motivation and Objective
Dataset
Methodology
Results & discussion
Conclusion
Outlook

## Slide 41
Three – feature dataset |  | 
Metrics | bce | bfoce
Seizure Sensitivity | 0.8 | 0.8
FA/24h | 453 | 441
Specificity | 0.8057 | 0.7726
Window Sensitivity | 0.6 | 0.54
bce = binary cross entropy , bfoce = binary focal cross entropy, FA/24h = False alarms per 24 hour
Five – feature dataset |  | 
Metrics | bce | bfoce
Seizure Sensitivity | 0.9 | 0.85
FA/24h | 382 | 366
Specificity | 0.8085 | 0.8338
Window Sensitivity | 0.696 | 0.68
Baseline selection – Choice of the Loss Function
41

## Slide 42
Five-feature dataset |  | 
Metrics | bce | bce + th_mv
Seizure Sensitivity | 0.9 | 0.8
FA/24h | 382 | 299
Specificity | 0.8085 | 0.8836
Window Sensitivity | 0.696 | 0.66
Baseline selection – Threshold Moving
bce = binary cross entropy, th_mv = threshold moving
42

## Slide 43
Three-feature dataset |  | 
Metrics | bce | bce + th_mv
Seizure Sensitivity | 0.8 | 0.8
FA/24h | 453 | 469
Specificity | 0.8057 | 0.791
Window Sensitivity | 0.6 | 0.63
Baseline selection – Threshold Moving
bce = binary cross entropy, th_mv = threshold moving
43

## Slide 44
Standard techniques – Five-feature dataset
Technique | FA/24h | Sensitivity seizure
Baseline | 382 | 0.9
Jittering (J) | 181 | 0.8
Rotation (R) | 351 | 0.9
Permutation (P) | 317 | 0.9
Time Warp (T) | 358 | 0.9
Magnitude Warp (M) | 242 | 0.85
Window Slicing (W) | 363 | 0.9
Perm + Time Warp (PT) | 343 | 0.9
Rot + Time Warp (RT) | 337 | 0.9
Rot + Perm (RP) | 259 | 0.9
Rot + Perm + Time Warp (RPT) | 298 | 0.85
44

## Slide 45
Technique | FA/24h | Sensitivity seizure
Baseline | 453 | 0.8
Jittering (J) | 228 | 0.7
Rotation (R) | 293 | 0.85
Permutation (P) | 422 | 0.85
Time Warp (T) | 551 | 0.95
Magnitude Warp (M) | 213 | 0.8
Window Slicing (W) | 540 | 0.85
Perm + Time Warp (PT) | 553 | 0.85
Rot + Time Warp (RT) | 297 | 0.85
Rot + Perm (RP) | 337 | 0.85
Rot + Perm + Time Warp (RPT) | 345 | 0.8
Standard techniques – Three-feature dataset
45

## Slide 46
Statement | Three-feature | Five-feature
FA/24h < baseline FA/24h | 7 | All techniques
Sensitivity > baseline Sensitivity | 7 | No technique
Sensitivity = baseline Sensitivity | 2 | 7
Sensitivity < baseline Sensitivity | 1 | 3
Better overall performance* | 6 | 7
Proportions | 4x, 8x | 1x, 2x, 4x
*Higher or same as baseline seizure sensitivity & lower than baseline FA/24h
Comparison of results for Standard techniques
46

## Slide 47
EpilepsyGAN |  | 
Model | FA/24h | Seizure Sensitivity
Baseline | 453 | 0.8
0% Overlap | 158 | 0.65
50% Overlap | 151 | 0.5
70% Overlap | 135 | 0.65
TTS-GAN |  | 
Model | FA/24h | Seizure Sensitivity
Baseline | 453 | 0.8
0% Overlap | 207 | 0.7
50% Overlap | 156 | 0.6
70% Overlap | 105 | 0.45
Best EpilepsyGAN model – data with 70% overlap
GANs – Selection of Overlap
Best TTS-GAN model – data with 0% overlap
Best models are used to generate synthetic data in 10%, 30%, 50%, 80% proportions for CNN training
47

## Slide 48
Model | FA/24h | Sensitivity Seizure
Baseline | 453 | 0.8
EpilepsyGAN (30%) | 330 | 0.8
TTS-GAN CS (50%) | 336 | 0.8
EpilepsyGAN (10%) | 416 | 0.85
TTS-GAN MSE (10%) | 414 | 0.85
Only a small proportion of synthetic data yields good results
GANs – Comparison of the results
48

## Slide 49
EpilepsyGAN
TTS-GAN
GANs – PCA & t-SNE distribution plots
49

## Slide 50
Five-feature dataset |  |  | 
Metrics | Baseline | Random Search | Hyperband
Seizure Sensitivity | 0.9 | 0.9 | 0.9
FA/24h | 382 | 631 | 461
Three-feature dataset |  |  | 
Metrics | Baseline | Random Search | Hyperband
Seizure Sensitivity | 0.8 | 0.95 | 0.9
FA/24h | 453 | 575 | 579
Additional Experiments – Hyperparameter tuning
Five-feature – Seizure sensitivity did not improve, while the False Alarm rate increased
Hyperparameter tuning with two different tuning algorithms
Three-feature – Seizure sensitivity improved at the cost of increased False Alarm rate
Random Search – achieves the highest Seizure sensitivity
50

## Slide 51
Five-feature dataset |  |  | 
Ratio | 1:1 | 1:5 | 1:10
Seizure Sensitivity | 0.8 | 0.5 | 0.4
FA/24h | 453 | 84 | 47
Window Sensitivity | 0.696 | 0.352 | 0.282
Specificity | 0.8085 | 0.9733 | 0.9874
Undersampling with increasing class ratios
Seizure sensitivity & False Alarm rate decreased significantly
Additional Experiments – Undersampling with different ratios
51

## Slide 52: Content
Motivation and Objective
Dataset
Methodology
Results & discussion
Conclusion
Outlook

## Slide 53
Conclusion
The use of binary cross-entropy loss demonstrated better overall performance than using binary focal cross-entropy loss.
Threshold moving showed no better results; it either decreased the seizure sensitivity or caused an increase in the false alarm rate.

## Slide 54
Conclusion
Rotation
Permutation
Time Warping
Window Slicing
Permutation + Time Warping
Rotation + Time Warping
Rotation + Permutation
In the Five- feature dataset case, the following Standard Augmentation techniques achieved the same Seizure Sensitivity with a reduced False Alarm rate:
54

## Slide 55
Conclusion
Rotation
Permutation
Time Warping
Window Slicing
Permutation + Time Warping
Rotation + Time Warping
Rotation + Permutation
In the Five- feature dataset case, the following Standard Augmentation techniques achieved the same Seizure Sensitivity with a reduced False Alarm rate:
55
FA/24h: 385  259
Seizure sensitivity: 90% (18 of 20 seizure detected)
Best Performance

## Slide 56: Conclusion
Rotation
Permutation
Magnitude Warping
Rotation + Time Warping
Rotation + Permutation
Rotation + Permutation + Time Warping
In the Three-feature dataset case, the following standard augmentation techniques achieved the same or higher Seizure Sensitivity with a reduced False Alarm rate:

## Slide 57: Conclusion
Rotation
Permutation
Magnitude Warping
Rotation + Time Warping
Rotation + Permutation
Rotation + Permutation + Time Warping
In the Three-feature dataset case, the following standard augmentation techniques achieved the same or higher Seizure Sensitivity with a reduced False Alarm rate:
FA/24h: 453  293
Seizure sensitivity: 80% (16 of 20 seizure detected)  85% (17 of 20 seizure detected)
FA/24h: 453  213
Seizure sensitivity: 80% (16 of 20 seizure detected)

## Slide 58: Conclusion
  EpilepsyGAN (30%)
GAN-generated synthetic data did not outperform those of the standard techniques
Utilizing a small proportion of synthetic data from GANs improved results by achieving the same/higher Seizure Sensitivity with a reduced False Alarm rate:
FA/24h: 453  330
Seizure sensitivity: 80% (16 of 20 seizures detected)
FA/24h: 453  414
Seizure sensitivity: 80% (16 of 20 seizures detected)  85% (17 of 20 seizures detected)
  TTS-GAN MSE (10%)

## Slide 59: Additional findings
In the five-feature case: Hyperparameter tuning with Random Search and Hyperband
   did not improve Seizure sensitivity, increased the False alarm rate
In the three-feature case: Hyperparameter tuning with Random Search and Hyperband
   increased Seizure sensitivity, increased the False alarm rate
Best-performing - Random Search in a three-feature case
  80%
  453
Undersampling with different class ratio results in very low seizure detection
 95% seizure sensitivity
 575 FA/24h

## Slide 60
60
Limited data, time and computational recourses to explore GANs to the full extent
  Large number of hyperparameters & Hyperparameter tuning needs a very long training time
  GANs training requires more real seizure data
Different sampling frequencies of the acc data, heart rate and temperature
  Adapting GAN architecture to multiple inputs with different frequencies
  Applying some Standard Augmentation techniques results in unrealistic values for heart rate & temperature
Lack of research regarding similarity metrics to evaluate time series data
Limitations & Challenges

## Slide 61: Content
Motivation and Objective
Dataset
Methodology
Results
Conclusion
Outlook

## Slide 62
Outlook
Accurate seizure detection is critical for reducing risks in epilepsy patients, while a low false alarm rate is desirable for practical use.
 Diagnosis & Alarm system
  In three-feature, Time Warping achieves the highest seizure sensitivity  95%; 551FA/24h
  Comparing Rotation and Magnitude Warping in three-feature case :
    Rotation - one more seizure  17 of 20 seizures detected; 293 FA/24h
    Magnitude Warping - 80 fewer false alarms  16 of 20 seizures detected; 213 FA/24h

## Slide 63
Outlook
Additional research could focus on optimizing the hyperparameters of GANs
Other time series GANs may lead to higher-quality synthetic seizure data
Evaluating the effect of augmentation on a bigger test set
Investigating the relationship between the type of seizure and the model’s performance via extensive test set = model’s effectiveness in detecting specific types of seizures
Exploring if augmentation influences a particular type of seizure detection
  related to sample pattern
Data quality

## Slide 64
Thank you For Your Attention!

## Slide 65: References
Image Sources:
WATSON, JL, and Lee Memorial. “Epilepsy Monitoring Unit Helps Physicians Make Diagnosis While Patients Sleep.” Fort Myers Florida Weekly -, Fort Myers Florida Weekly, 10 Sept. 2014, fortmyers.floridaweekly.com/articles/epilepsy-monitoring-unit-helps-physicians-make-diagnosis-while-patients-sleep/. Accessed 30 May 2023.
Houta, Salima, Pinar Bisgin, and Pascal Dulich. "Machine learning methods for detection of Epileptic seizures with long-term wearable devices." Eleventh International Conference on eHealth, Telemedicine, and Social Medicine. 2019.
“MOND -Mobile, Smart Neurosensor System for the Detection and Documentation of Epileptic Seizures in Everyday Life.” Fraunhofer Institute for Software and Systems Engineering, 2020, www.isst.fraunhofer.de/en/business-units/healthcare/projects/MOND.html. Accessed 30 May 2023.
“Papers with Code - DTW Explained.” Paperswithcode.com, 2014, paperswithcode.com/method/dtw. Accessed 30 May 2023.
Rosebrock, Adrian. “Building an Image Search Engine: Defining Your Similarity Metric (Step 3 of 4) - PyImageSearch.” PyImageSearch, 17 Feb. 2014, pyimagesearch.com/2014/02/17/building-an-image-search-engine-defining-your-similarity-metric-step-3-of-4/. Accessed 30 May 2023.
Gupta, Aryan. “Mean Squared Error : Overview, Examples, Concepts and More.” Simplilearn.com, Simplilearn, 15 Dec. 2021, www.simplilearn.com/tutorials/statistics-tutorial/mean-squared-error. Accessed 30 May 2023.
Convolutional Neural Networks:
Meisel, Christian et al. “Machine learning from wristband sensor data for wearable, noninvasive seizure forecasting.” Epilepsia vol. 61,12 (2020): 2653-2666. doi:10.1111/epi.16719
Böring, Jannik. "Implementierung eines Convolutional Neural Networks zur Detektion motorischer epileptischer Anfälle mittels Biosignalen eines In-Ohr-Sensors", 2021.
Generative Adversarial Networks:
Li, Xiaomin, et al. "Tts-gan: A transformer-based time-series generative adversarial network." Artificial Intelligence in Medicine: 20th International Conference on Artificial Intelligence in Medicine, AIME 2022, Halifax, NS, Canada, June 14–17, 2022, Proceedings. Cham: Springer International Publishing, 2022.
Pascual, Damian, et al. "Epilepsygan: Synthetic epileptic brain activities with privacy preservation." IEEE Transactions on Biomedical Engineering 68.8 (2020): 2435-2446.
Standard Augmentation Techniques
Iwana, Brian Kenji, and Seiichi Uchida. "An empirical survey of data augmentation for time series classification with neural networks." Plos one 16.7 (2021): e0254841.
Um, Terry T., et al. "Data augmentation of wearable sensor data for parkinson’s disease monitoring using convolutional neural networks." Proceedings of the 19th ACM international conference on multimodal interaction. 2017.

## Slide 66: References
“Epilepsy Foundation.” Epilepsy Foundation, 2023, www.epilepsy.com/. Accessed 30 May 2023.
“Welcome to the International League against Epilepsy.” Ilae.org, 2017, www.ilae.org/. Accessed 30 May 2023.
MOND PROJECT, https://www.isst.fraunhofer.de/de/geschaeftsfelder/gesundheitswesen/projekte/MOND.html. Accessed 23 March 2023.
Henze, Jasmin, et al. "Multimodal detection of tonic–clonic seizures based on 3D acceleration and heart rate data from an in-ear sensor." Pattern Recognition. ICPR International Workshops and Challenges: Virtual Event, January 10–15, 2021, Proceedings, Part I. Springer International Publishing, 2021.
Beniczky, Sándor, and Philippe Ryvlin. “Standards for testing and clinical validation of seizure detection devices.” Epilepsia vol. 59 Suppl 1 (2018): 9-13. doi:10.1111/epi.14049
Sun, Yanmin, Andrew KC Wong, and Mohamed S. Kamel. "Classification of imbalanced data: A review." International journal of pattern recognition and artificial intelligence 23.04 (2009): 687-719.
He, Haibo, and Edwardo A. Garcia. "Learning from imbalanced data." IEEE Transactions on knowledge and data engineering 21.9 (2009): 1263-1284.
Brownlee, Jason. “A Gentle Introduction to Threshold-Moving for Imbalanced Classification - MachineLearningMastery.com.” MachineLearningMastery.com, 9 Feb. 2020, machinelearningmastery.com/threshold-moving-for-imbalanced-classification/. Accessed 23 March 2023.
Lin, Tsung-Yi, et al. "Focal loss for dense object detection." Proceedings of the IEEE international conference on computer vision. 2017.
Brownlee, Jason. “A Gentle Introduction to Cross-Entropy for Machine Learning - MachineLearningMastery.com.” MachineLearningMastery.com, 20 Oct. 2019, machinelearningmastery.com/cross-entropy-for-machine-learning/. Accessed 6 May 2023.
Wold, Svante, Kim Esbensen, and Paul Geladi. "Principal component analysis." Chemometrics and intelligent laboratory systems 2.1-3 (1987): 37-52.
Van der Maaten, Laurens, and Geoffrey Hinton. "Visualizing data using t-SNE." Journal of machine learning research 9.11 (2008).
Bergstra, James, and Yoshua Bengio. "Random search for hyper-parameter optimization." Journal of machine learning research 13.2 (2012).
Li, Lisha, et al. "Efficient hyperparameter optimization and infinitely many armed bandits." CoRR, abs/1603.06560 16 (2016).

## Slide 68: Backup Slides

## Slide 69: Metrics

## Slide 70
  Seizure Sensitivity – over the whole seizure interval
  Window Sensitivity – over individual window
  Specificity
  False Alarms per 24 hours (FA/24h)
70
1
1
1
0
0
0
0
0
0
0
0
Seizure 1
Detected
Not Detected
1
0
0
0
0
0
0
1
0
0
Detected
Seizure 2
Seizure 3
Evaluation

## Slide 71
  Seizure Sensitivity – over the whole seizure interval
71
Evaluation
Detected
Not Detected
Detected
1
0
= 0.666
Seizure 1
Seizure 2
Seizure 3

## Slide 72
  Seizure Sensitivity – over the whole seizure interval
  Window Sensitivity – over individual window
72
Evaluation
Seizure 1
Detected
Not Detected
Detected
Seizure 2
Seizure 3
TP
TP
TP
TP
FN
FN
FN
FN
FN
FN
FN
FN
1
0
= 0.333

## Slide 73
  Seizure Sensitivity – over the whole seizure interval
  Window Sensitivity – over individual window
  Specificity
73
Evaluation
Seizure 1
Seizure 2
Seizure 3
1
0
TN
TN
TN
TN
TN
TN
FP
TN
FP
= 0.777

## Slide 74
  Seizure Sensitivity – over the whole seizure interval
  Window Sensitivity – over individual window
  Specificity
  False Alarms per 24 hours (FA/24h)
    Base on average seizure interval (approx. 90s)
74
Evaluation
Total Non-seizure duration in days

## Slide 75
  Seizure Sensitivity – over the whole seizure interval
  Window Sensitivity – over individual window
  Specificity
  False Alarms per 24 hours (FA/24h)
75
Evaluation
Time (t)
90 seconds interval
1 1 1 1 1 1 1 1 1
FP_90s
1 1 0 1 0 0 1
1 1 1 1 1 1
1 1 1 1 1 1 1 1
FP_90s
FA / FP_90s
No FA / TN_90s
Example
# of Non-seizure windows = 59370
FP_90s = 2624
Total Non-seizure duration in days =
FA/24h = 2624/6.87152 = 382

## Slide 76: Why FA/24h is lower while having lower specificity
Time (t)
90 seconds interval
0 1 1 1
0 1 1 1
1 1 1 0
0 1 1 1
No FA
No FA
No FA
No FA
Time (t)
1 1 1 1
0 0 0 0
0 0 0 0
1 1 1 1
FA
No FA
No FA
FA
4 non-seizure windows correctly classified, 0 FA/24h
8 non-seizure windows correctly classified, 2 FA/24h
Low specificity at low FA/24h
High specificity at High FA/24h

## Slide 77
Seizure Interval 1
Seizure Interval 2
Detected 10s Window
Non - Detected 10s Window
Model 1
Model 2
Seizure Interval 2
Seizure Interval 1
Comparison between Seizure sensitivity and Sample Sensitivity
Seizure detected = 1 of 2
Window detected = 3 of 7
Not detected
detected
detected
detected
Seizure detected = 2 of 2
Window detected = 2 of 7
Low seizure sensitivity
High Sample/window sensitivity
High seizure sensitivity
Low Sample/window sensitivity

## Slide 78: Hyperparameter

## Slide 79
79
Hyperparameter | Interval
No. of filters | 32, 64, 128, 256
Filter size | Min=2 , Max =10 , Step size = 1
Size of Pooling | Min=2 , Max =10 , Step size = 1
Dropout rate | Min=0.2 , Max =0.9 , Step size = 0.1
Number of Conv1D, Pooling, Dropout group | Min=1 , Max =4 , Step size = 1
Learning rate | 0.01, 0.001, 0.0001
Dense layer neurons | Min=20 , Max =100, Step size = 10
Hyperparameter Intervals

## Slide 80: G-mean

## Slide 81
Threshold Moving
Sensitivity = True Positive Rate
Specificity = 1- False Positive Rate
False Positive Rate = FP/FP+TN
  TPR and FPR for multiple threshold values are computed.
  The G-mean for each threshold value is calculated by: G-mean =
  The threshold that results in the maximum G-mean value is selected.

## Slide 82: GANs

## Slide 83
TTS-GAN - Network architecture
Based on Source: Li, Xiaomin, et al. "Tts-gan: A transformer-based time-series generative adversarial network." Artificial Intelligence in Medicine: 20th International Conference on Artificial Intelligence in Medicine, AIME 2022, Halifax, NS, Canada, June 14–17, 2022, Proceedings. Cham: Springer International Publishing, 2022.

## Slide 84
Latent Dimension: The latent dimension refers to a size of the noise vector consisting of random variables generated from a random normal (Gaussian) distribution. This noise vector serves as the input to the generator in a GAN architecture. In our case it is 500 -> (*, 500)
Embedding Dimension: The embedding dimension refers to the size of the vector space in which the input data is projected or embedded. It plays a crucial role in shaping the dimensionality of the learned representations and affects the model's capacity to capture and represent complex patterns in the data.
  In the context of transformers, the embedding dimension refers to the size of the vector space in which words (or in this case, patches of the time-series data) are embedded. E.g (*, 500) -> (*,seq_len x emb_dim) -> (*, 500 x 10) -> (*,500,10)
Patch Size: The patch size refers to the dimensions of the patches that are extracted from the input. This division allows the discriminator to analyze the time-series data in a localized manner.
  In the context of transformers applied to images or time-series data, the input is often divided into patches. Each patch is then processed independently by the transformer. This approach allows the model to handle larger inputs and learn local and global patterns.

## Slide 85
Positional Encoding: Soft positional encoding values are added to the end of each patch. These positional values are learned during model training and help the model differentiate between different positions within each patch. The positional encoding ensures that the model can capture temporal relationships and dependencies between the timesteps

## Slide 86
TTS-GAN - Workflow
Seizures
Original Test set
Seizures
Original Train set
Training
TTS-GAN
Trained TTS - GAN
Synthetic Seizures Data
Synthetic Seizures Data
Similarity Evaluation
Original Seizure & non-Seizure
CNN-Training
Evaluation
DATASET
Random noise

## Slide 87
Overlapping | # of Seizure windows
0% | 585
50% | 1113
70% | 1638
TTS-GAN – Training
Activation function: GeLU
Optimizer: Adam
Batch size: 8; epochs: 500
Lr generator: 0.0001 ; Lr discriminator: 0.0003
Loss function: MSE (Mean Squared Error)
Hyperparameters:
  Embedding dimension: 10
  Latent dimension: 100
  Patch size: 50
Source: Li, Xiaomin, et al. "Tts-gan: A transformer-based time-series generative adversarial network." Artificial Intelligence in Medicine: 20th International Conference on Artificial Intelligence in Medicine, AIME 2022, Halifax, NS, Canada, June 14–17, 2022, Proceedings. Cham: Springer International Publishing, 2022.

## Slide 88
TTS-GAN

## Slide 89
Epilepsy GAN - Network architecture
Based on source: Pascual, Damian, et al. "Epilepsygan: Synthetic epileptic brain activities with privacy preservation." IEEE Transactions on Biomedical Engineering 68.8 (2020): 2435-2446.

## Slide 90
Patient 1
Patient i-1
Patient i
Patient i+1
…
…
Patient N
Non-Seizures
Seizures
GAN
Training set GAN
Real
Training
Synthetic
Real
Seizure
Non- Seizure
Seizures
CNN Model Training

## Slide 91
Epilepsy GAN - Workflow
Seizures
Non-Seizures
Dataset
Test Patients Data
Seizures
Non-Seizures
Training Patients Data
Training Epilepsy GAN
A
B
Synthetic Seizures
Original Seizure & non-Seizure
CNN-Training
Pairing
Seizures
Similarity Evaluation
Trained Epilepsy GAN
Trained Epilepsy GAN
Synthetic Seizures

## Slide 92
EpilepsyGAN

## Slide 93
Overlapping | # of Seizure windows | # of Seizure & Non-seizure pairs
0% | 525 | 5250
50% | 1009 | 10090
70% | 1484 | 14840
Activation function: leaky ReLu
Optimizer: Adam
Batch size: 100; epochs: 100
Lr generator: 0.0001 ; Lr discriminator: 0.0004
Loss function: LSGAN (Least Squares GAN)
Regularization: L1 , λ: 100
Epilepsy GAN – Training

## Slide 94: CNN

## Slide 95
CNN architecture – illustration of 1D CNN
Based on Source: Böring, Jannik. "Implementierung eines Convolutional Neural Networks zur Detektion motorischer epileptischer Anfälle mittels Biosignalen eines In-Ohr-Sensors", 2021.

## Slide 96
If window fails, next window starts after 1 second.

## Slide 97
Training of CNN
Activation function: ReLU, Softmax
Optimizer: Adam
Batch size: 2048; epochs: 2000
Learning rate: scheduled with Learning rate scheduler function from Keras API
Loss function: binary cross-entropy, binary-focal cross-entropy
Dropout: 0.7 is used for regularization

## Slide 98
Training of CNN
Activation function: ReLU, Softmax
Optimizer: Adam
Batch size: 32; epochs: 2000
Learning rate: scheduled with Learning rate scheduler function from Keras API
Loss function: binary cross-entropy, binary-focal cross-entropy
Dropout: 0.7 is used for regularization
Data distribution in the five-feature dataset
  84% train: 660 seizure & non-seizure windows
  16% test: 125 seizure windows
Data distribution in a three-feature dataset
  84% train: 791 seizure & non-seizure windows
  16% test: 150 seizure windows

## Slide 99: SIMILARITY

## Slide 100
Dynamic Time Warping (DTW)
  Measures the similarity between two-time series signals by finding the optimal alignment of the two signals such that the difference between them is minimized
Similarity Evaluation
Source: “Papers with Code - DTW Explained.” Paperswithcode.com, 2014, paperswithcode.com/method/dtw. Accessed 30 May 2023.
‌

## Slide 101
Dynamic Time Warping (DTW) – Calculation
  Distance Matrix – Distance between each corresponding point from two time series
Euclidean distance 
Similarity Evaluation
Source: Shou, Yutao, Nikos Mamoulis, and David W. Cheung. "Fast and exact warping of time
series using adaptive segmental approximations." Machine Learning 58.2-3 (2005): 231.

## Slide 102
Dynamic Time Warping (DTW) – Calculation
  Distance Matrix – Distance between each corresponding point from two time series
  Warping Matrix /Cost Matrix
Similarity Evaluation
Source: Shou, Yutao, Nikos Mamoulis, and David W. Cheung. "Fast and exact warping of time
series using adaptive segmental approximations." Machine Learning 58.2-3 (2005): 231.

## Slide 103
Dynamic Time Warping (DTW) – Calculation
  Distance Matrix – Distance between each corresponding point from two time series
  Warping Matrix – Cost spent
  DTW is calculated for each pair of real and synthetic samples and then the average is taken
Similarity Evaluation

## Slide 104
Dynamic Time Warping (DTW) – Calculation
  Distance Matrix – Distance between each corresponding point from two time series
  Warping Matrix – Cost spent
  DTW is calculated for each pair of real and synthetic samples and then the average is taken
  Values between 0 – max. real DTW
Similarity Evaluation

## Slide 105
Dynamic Time Warping (DTW)
Cosine Similarity
  Calculated between the feature vectors of two time series and measures the similarity by calculating the cosine of the angle between them
Similarity Evaluation
Source: Rosebrock, Adrian. “Building an Image Search Engine: Defining Your Similarity Metric (Step 3 of 4) - PyImageSearch.” PyImageSearch, 17 Feb. 2014, pyimagesearch.com/2014/02/17/building-an-image-search-engine-defining-your-similarity-metric-step-3-of-4/. Accessed 30 May 2023.
‌

## Slide 106
Dynamic Time Warping (DTW)
Cosine Similarity – Calculation
  Seven statistical features are calculated for each real and synthetic sample;
    median, mean, standard deviation, variance, root mean square, maximum, and minimum
Similarity Evaluation

## Slide 107
Dynamic Time Warping (DTW)
Cosine Similarity – Calculation
  Seven statistical features are calculated for each real and synthetic sample;
    median, mean, standard deviation, variance, root mean square, maximum, and minimum
  Cosine similarity is calculated for each pair of real and synthetic feature vectors then the average is taken
Similarity Evaluation

## Slide 108
Dynamic Time Warping (DTW)
Cosine Similarity – Calculation
  Seven statistical features are calculated for each real and synthetic sample;
    median, mean, standard deviation, variance, root mean square, maximum, and minimum
  Cosine similarity is calculated for each pair of real and synthetic feature vectors then the average is taken
  Values between -1 and 1
Similarity Evaluation

## Slide 109
Dynamic Time Warping (DTW)
Cosine Similarity
Mean Squared Error
  Assesses the difference between the real and synthetic values / (value-wise comparison) of a time series at each time point. The lower MSE value refers to the higher similarity between the two-time series
Similarity Evaluation
Source: Gupta, Aryan. “Mean Squared Error : Overview, Examples, Concepts and More.” Simplilearn.com, Simplilearn, 15 Dec. 2021, www.simplilearn.com/tutorials/statistics-tutorial/mean-squared-error. Accessed 30 May 2023.
‌

## Slide 110
Dynamic Time Warping (DTW)
Cosine Similarity
Mean Squared Error – Calculation
  Calculated for each pair of real and synthetic samples and then the average is taken
Similarity Evaluation

## Slide 111
Dynamic Time Warping (DTW)
Cosine Similarity
Mean Squared Error – Calculation
  Calculated for each pair of real and synthetic samples and then the average is taken
  Values between 0 and 1
Similarity Evaluation

## Slide 112
Dynamic Time Warping (DTW)
Cosine Similarity
Mean Squared Error
PCA and t-SNE
  Visualizing the distribution of higher dimensional data in a two-dimensional space
Similarity Evaluation

## Slide 114: EXTRA

## Slide 115
Linear Transform
Noise 1D Input Vector
Patch Division
Positional Embedding
Layer Norm
Layer Norm
Multi Head Attention
Feed Forward
Dropout
Dropout
Repeat
N x
Conv2D channel reduction
Synthetic Signals
Generator

## Slide 116
Layer Norm
Layer Norm
Multi Head Attention
Feed Forward
Dropout
Dropout
Repeat
M x
Synthetic Signals & Real Signals
Discriminator
Patch Embedding Linear
Classification Head
Synthetic / Real Signals ?

## Slide 117
Statement | Three-feature | Five-feature
FA/24h < baseline FA/24h | 7 | All techniques
Sensitivity > baseline Sensitivity | 7 | No technique
Sensitivity = baseline Sensitivity | 2 | 7
Sensitivity < baseline Sensitivity | 1 | 3
Better overall performance* | 6 | 7
Proportions | 4x, 8x | 1x, 2x, 4x
*Higher or same as baseline seizure sensitivity & lower than baseline FA/24h
Except Time Warping, no technique with a three-feature dataset achieved higher seizure sensitivity than the same technique in five-feature dataset
Time Warping with three-feature dataset yields the higher seizure sensitivity (0.95) than it does in five-feature case, at the cost of a much higher false alarm rate
Rotation, Rotation + Time Warping, and Magnitude Warping result in lower FA/24h in three-feature case, while all others give lower FA/24h in five-feature case
Comparison of results for Standard techniques
117

## Slide 118
Window sensitivity vs. Specificity in five-feature dataset

## Slide 119
Window sensitivity vs. Specificity in three-feature dataset

## Slide 120
Loss Curves 5-feat
BCE
BFOCE

## Slide 121
BCE
BFOCE
Loss Curves 3-feat

## Slide 122
Even though with the use of data augmentation the seizure sensitivity is higher and the
false alarm rate is lower as compared to the baseline model (model results without augmentation),
the model is still not suitable for usage in practice. The false alarm rate is still
too high; 12 alarms per hour (R) and 9 alarms per hour (M) for the model trained along
with data from the best-performing augmentation techniques in three-feature case. In the
five-feature case, the best-performing technique (RP) yields 11 false alarms per hour.

## Slide 123
14
50
10
11
13
40
2
1
4
118
154
83
195
Filter
Input
Output
2*50 + 1 * 10 + 4*11